MR660 PC Application Overview

By: G. Chopcinski, Kg Solutions, LLC

Date: Mar 19, 2022 Supports V1.5.4

## Splash screen

|  |  |
| --- | --- |
|  | This screen is displayed when powering up the MR660. |

## Standby screen

|  |  |
| --- | --- |
|  | This screen is displayed when the MR660 has not detected a Mouthpiece aperture operation or the Play/Pause button has not been used within the last 30 minutes.   * No changes needed. * [gchop] the logo is larger. |

## Mouthpiece processing screens

|  |  |
| --- | --- |
|  | This screen instructs the user to Insert the Mouthpiece.   * Bold and Italicize “Insert Mouthpiece”. See slide 2 in flowchart. * [chop] Changed as shown. |
|  | This screen indicates that the Mouthpiece (electronics) is being read by the MR660.  This screen should be displayed briefly. |
|  | This screen indicates that the Mouthpiece has been read correctly and is prompting the User to verify the correct Serial Number.  If this screen is displayed then the Mouthpiece information from the electronics (EEPROM) is valid AND the Mouthpiece has not been used in a completed Therapy within the last 12 hours AND the number of Mouthpieces uses is less than the Life-Usage.   * Bold and Italicize “Press Play/Pause to Continue”. See slide 2 in flowchart. * [chop] changed as shown. |
|  | This screen is displayed when the User has accepted the above-mentioned Mouthpiece Serial Number screen and the Patient is ready to start Therapy.   * Bold and Italicize “Press Play/Pause to Start Therapy”. See slide 2 in flowchart. * [chop] Changed as shown. |

|  |  |
| --- | --- |
|  | This Display appears after machine shuts down during operation and the patient is going through the steps to resume therapy again.   * Bold and Italicize “Press Play/Pause to Resume Therapy”. See slide 2 in flowchart. * [chop] Changed as shown. |

## Therapy screens

|  |  |  |
| --- | --- | --- |
|  | | These screens are displayed during Therapy.  The green and white ticks depict the remaining seconds and the number indicates the number of minutes remaining of the Therapy. |
|  | |  |
|  | | Indicates Therapy completion. |
|  | This screen is displayed when the Play/Pause button is pressed during Therapy. | |

## Error Conditions

|  |  |
| --- | --- |
|  | This screen is displayed if the Mouthpiece electronics cannot be read properly.  This can occur while first processing the Mouthpiece or during Therapy while attempting to update the information in the Mouthpiece.   * Bold and Italicize “Reinsert Mouthpiece” on this document only. * During GUI Operation when this screen is displayed “Reinsert Mouthpiece” doesn’t disappear when you detach the mouthpiece, this needs to be fixed. See slide 8 (MCA is Removed during operation) in the flowchart.  [CHOP] It does follow Slide 8. Please confirm. |
|  | This screen is displayed if the Mouthpiece has been used too many times.   * Keep the wording “Mouthpiece Expired” on v1.5.3 it says, “Mouthpiece Daily Limit Reached” and this is not correct. See slide 6 in flowchart. |
|  | This screen is displayed if the Mouthpiece successfully completed a Therapy within the last 12 hours.   * Change to “Mouthpiece Daily Limit Reached”. See slide 9 in the flowchart. |
| Logo  Description automatically generated | This screen is displayed whenever a system error has been detected. The number “01” is replaced with the appropriate Error Code.  This screen is displayed when mouthpiece is detached during operation   * Should say “Mouthpiece Detached Reinsert Mouthpiece” * Bold and Italicize “Reinsert Mouthpiece” * See slide 8 in the flowchart. |