## normal states
Software Normal State Screen Messages
# | Image | Text Color | Background Color | Wording | LED Ring Color | LED Ring Flashing?
1 | new | White | Black | INSERT MOUTHPIECE | White | no
2 | new | White | Black | READING... | Blue | no
3 | new | White | Black | S/N MCA0x#### Press PLAY if correct | Blue | no
4 | new | White | Black | READY Press PLAY | Blue | no
5 | new | White | Black | countdown clock images | Green | no
6 | new | White | Black | THERAPY COMPLETE | Green | yes

## error states
Software Error State Screen Messages
# | Image | Text Color | Background Color | Wording | LED Ring Color | LED Ring Flashing? | Purpose
1 | new | White | Black | CONTROL UNIT ERROR 01 | red | no | Alarm Error
2 | new | White | Black | CONTROL UNIT ERROR 02 | red | no | Current Level Error
3 | new | White | Black | CONTROL UNIT ERROR 03 | red | no | Current src Error
4 | new | White | Black | CONTROL UNIT ERROR 04 | red | no | Current too high
5 | new | White | Black | CONTROL UNIT ERROR 05 | red | no | Current too low
6 | new | White | Black | CONTROL UNIT ERROR 06 | red | no | Error 5VDC
7 | new | White | Black | CONTROL UNIT ERROR 07 | red | no | Error 3.3 VDC
8 | new | White | Black | CONTROL UNIT ERROR 08 | red | no | Error 24 VDC
9 | new | White | Black | CONTROL UNIT ERROR 09 | red | no | Start wdog error
10 | new | White | Black | CONTROL UNIT ERROR 10 | red | no | Error therapy on
11 | new | White | Black | CONTROL UNIT ERROR 11 | red | no | Watchdog enable error
12 | new | White | Black | CONTROL UNIT ERROR 12 | red | no | Fan error
13 | new | White | Black | CONTROL UNIT ERROR 13 | red | no | Flash mem error
14 | new | White | Black | CONTROL UNIT ERROR 14 | red | no | HW wdog timeout
15 | new | White | Black | CONTROL UNIT ERROR 15 | red | no | I2C error
16 | new | White | Black | CONTROL UNIT ERROR 16 | red | no | LCD error
17 | new | White | Black | CONTROL UNIT ERROR 17 | red | no | EEPROM error
18 | new | White | Black | CONTROL UNIT ERROR 18 | red | no | Button Error
19 | new | White | Black | CONTROL UNIT ERROR 19 | red | no | RGB LED error
20 | new | White | Black | CONTROL UNIT ERROR 20 | red | no | SW wdog timeout
21 | new | White | Black | CONTROL UNIT ERROR 21 | red | no | Temp too high
22 | new | White | Black | CONTROL UNIT ERROR 22 | red | no | Unknown error
23 | new | White | Black | MOUTHPIECE DETACHED | red | no | Handle detached, limit switch not acutated
24 | new | White | Black | MOUTHPIECE EXPIRED | red | no | Total time of MCA use is >12,300 seconds
25 | new | White | Black | MOUTHPIECE READING ERROR | red | no | Either a connection problem or it fails to write to EEPROM 5 times

## other screens
Other Screens
# | Image | Text Color | Background Color | Wording | LED Ring Color | LED Ring Flashing? | Purpose | Details
1 | see below | N/A | White | N/A, it’s the MuReva Logo | White | no | Start up splash screen
2 | see below | N/A | Black | N/A, it's the second MuReva Logo | LED ring color is whatever color it was prior to entering standby state | no | Screensaver | starts after 30 minutes of inactivity
runs for 30 minutes and then transitions to blank/LCD is turned off
Splash picture
Screensaver:

## Old reference pics
# | Image | Purpose
1
2
3
4
1 |  | Alarm Error
2 |  | Current Level Error
3 |  | Current src Error
4 |  | Current too high
5 |  | Current too low
6 |  | Error 5VDC
7 |  | Error 3.3 VDC
8 |  | Error 24 VDC
9 |  | Start wdog error
10 |  | Error therapy on
11 |  | Watchdog enable error
12 |  | Fan error
13 |  | Flash mem error
14 |  | HW wdog timeout
15 |  | I2C error
16 |  | LCD error
17 |  | EEPROM error
18 |  | Button Error
19 |  | RGB LED error
20 |  | SW wdog timeout
21 |  | Temp too high
22 |  | Unknown error
23 |  | Handle detached, limit switch not acutated