## Slide 1
Start
Power Up (MuReva Logo splash screen)
INSERT MOUTHPIECE
READING...
READY
Press PLAY
PAUSED
Operation (countdown clock from 5 min)
THERAPY COMPLETE
MOUTHPIECEREADING ERROR
CONTROL UNIT
ERROR ##
S/N MCA0x#### Press PLAY if correct
MOUTHPIECE EXPIRED
MOUTHPIECE DETACHED
Power Switch
Flipped On
5 seconds elapsed
MCA
inserted
Play/Pause pressed
Play/Pause pressed
Play/Pause pressed
MCA
Removed
MCA
Removed
MCA
Removed
MCA
Removed
MCA
inserted
MCA already
inserted
Same MCA
inserted
Standby Mode (bouncing logo screen)
No user action after 30 min
No user action after 30 min
No user action after 1 hour (screen goes to blank)
Blinking green ring

## Slide 2
Start
Power Up (MuReva Logo splash screen)
INSERT MOUTHPIECE
READING...
READY
Press PLAY
PAUSED
Operation (countdown clock from 5 min)
THERAPY COMPLETE
S/N MCA0x#### Press PLAY if correct
Power Switch
Flipped On
5 seconds elapsed
MCA
inserted
Play/Pause pressed
Play/Pause pressed
Play/Pause pressed
MCA
Removed
MCA already
inserted
Normal Operation
Blinking green ring

## Slide 3
Start
Power Up (MuReva Logo splash screen)
INSERT MOUTHPIECE
READING...
READY
Press PLAY
PAUSED
Operation (countdown clock from 5 min)
THERAPY COMPLETE
CONTROL UNIT
ERROR ##
S/N MCA0x#### Press PLAY if correct
Power Switch
Flipped On
5 seconds elapsed
MCA
inserted
Play/Pause pressed
Play/Pause pressed
Play/Pause pressed
Standby Mode (bouncing logo screen)
Control Unit Error Path
- ## is the Control Unit error number
- LCU must be restarted via on/off switch to clear error code (LCU does NOT enter standby)
Blinking green ring

## Slide 4
Start
Power Up (MuReva Logo splash screen)
INSERT MOUTHPIECE
READING...
READY
Press PLAY
PAUSED
Operation (countdown clock from 5 min)
THERAPY COMPLETE
MOUTHPIECEREADING ERROR
S/N MCA0x#### Press PLAY if correct
Power Switch
Flipped On
5 seconds elapsed
MCA
inserted
Play/Pause pressed
Play/Pause pressed
Play/Pause pressed
MCA
Removed
MCA
Removed
MCA already
inserted
Mouthpiece Reading Error Occurs During “Reading...” phase
-remove MCA to return to “insert mouthpiece” screen
Blinking green ring

## Slide 5
Start
Power Up (MuReva Logo splash screen)
INSERT MOUTHPIECE
READING...
READY
Press PLAY
PAUSED
Operation (countdown clock from 5 min)
THERAPY COMPLETE
MOUTHPIECEREADING ERROR
S/N MCA0x#### Press PLAY if correct
Power Switch
Flipped On
5 seconds elapsed
MCA
inserted
Play/Pause pressed
Play/Pause pressed
Play/Pause pressed
MCA
Removed
MCA
Removed
MCA already
inserted
Same MCA
inserted
Mouthpiece Reading Error Occurs During Operation phase
-Cause: missed writing to EEPROM on MCA 5 times in a row
-remove MCA to return to “insert mouthpiece” screen
Blinking green ring

## Slide 6
Start
Power Up (MuReva Logo splash screen)
INSERT MOUTHPIECE
READING...
READY
Press PLAY
PAUSED
Operation (countdown clock from 5 min)
THERAPY COMPLETE
S/N MCA0x#### Press PLAY if correct
MOUTHPIECE EXPIRED
Power Switch
Flipped On
5 seconds elapsed
MCA
inserted
Play/Pause pressed
Play/Pause pressed
Play/Pause pressed
MCA
Removed
MCA
Removed
MCA already
inserted
Mouthpiece Reading determines that the MCA is expired
-remove device to return to “insert mouthpiece” screen
Blinking green ring

## Slide 7
Start
Power Up (MuReva Logo splash screen)
INSERT MOUTHPIECE
READING...
READY
Press PLAY
PAUSED
Operation (countdown clock from 5 min)
THERAPY COMPLETE
S/N MCA0x#### Press PLAY if correct
MOUTHPIECE DETACHED
Power Switch
Flipped On
5 seconds elapsed
MCA
inserted
Play/Pause pressed
Play/Pause pressed
Play/Pause pressed
MCA
Removed
MCA
Removed
MCA
inserted
MCA already
inserted
MCA is removed in “S/N MCA0x#### Press PLAY if correct” OR “READY Press PLAY”
-insert mouthpiece to initiate “READING...” screen and re-reading process
Blinking green ring

## Slide 8
Start
Power Up (MuReva Logo splash screen)
INSERT MOUTHPIECE
READING...
READY
Press PLAY
PAUSED
Operation (countdown clock from 5 min)
THERAPY COMPLETE
S/N MCA0x#### Press PLAY if correct
MOUTHPIECE DETACHED
Power Switch
Flipped On
5 seconds elapsed
MCA
inserted
Play/Pause pressed
Play/Pause pressed
Play/Pause pressed
MCA
Removed
MCA
Removed
MCA
inserted
MCA already
inserted
Same MCA
inserted
MCA is removed during operation
-reinsert mouthpiece to go to “PAUSED” screen
-press play to resume normal operation
Blinking green ring

## Slide 9
Start
Power Up (MuReva Logo splash screen)
INSERT MOUTHPIECE
READING...
READY
Press PLAY
PAUSED
Operation (countdown clock from 5 min)
THERAPY COMPLETE
MOUTHPIECEREADING ERROR
S/N MCA0x#### Press PLAY if correct
MOUTHPIECE EXPIRED
MOUTHPIECE DETACHED
Power Switch
Flipped On
5 seconds elapsed
MCA
inserted
Play/Pause pressed
Play/Pause pressed
Play/Pause pressed
MCA already
inserted
Standby Mode (bouncing logo screen)
No user action after 30 min
No user action after 30 min
No user action after 1 hour (screen goes to blank)
Standby Mode
-bouncing logo screen
-occurs when there is no user action:
-user does not insert MCA during “INSERT MOUTHPIECE”, “MOUTHPIECE DETACHED”, or “PAUSED”
-inserting a MCA wakes up the control unit and the control unit goes to “READING...”
-user does not remove the MCA during “THERAPY COMPLETE”, “MOUTHPIECE EXPIRED”, “MOUTHPIECE READING ERROR”
-removing the MCA wakes up the control unit and the control unit goes to “INSERT MOUTHPIECE”
-user does not press pushbutton during “S/N MCA0x####”, “READY Press PLAY”, or “PAUSED”
-pressing the pushbutton wakes up the control unit, then press play again
-after 1 hour total has elapsed, screen goes to blank and user can only restart the control unit via power on/off switch
Blinking green ring

## Slide 10
Other Requirements:
If power is cut off via power switch/power cord, go back to start
Play button is only functional during the “SN MCAXXXXXX Press Play if Correct”, “READY Press PLAY” or “PAUSED” states
Pause is only functional during the operation state
If therapy timer has started, pressing the Play button or if the MCA disconnects will pause the therapy and the countdown
LED therapy lights only turn on when Play button is pressed. They are turned off during pauses, disconnects, why therapy is complete, etc
Standby LED ring color is any color it was prior to entering standby
To get out of Standby mode, press play/pause pushbutton
During the “READING...” phase, the control unit is reading the S/N, calibration number, and that the total time is less than 12,300 seconds